## Slide 1: Mechanical Circulatory Assist Devices April 20, 2023

## Slide 2
Discuss available mechanical circulatory support devices for home use
Review clinical management for rehabilitation of patients with mechanical circulatory support devices
To understand patient care requirements during emergency situations
Learning Objectives
Notes: 2
The importance of ventricular assist devices with cardiovascular disease

## Slide 3
Heart disease is the leading cause of death for men, women, and people of most racial and ethnic groups in the United States.1
One person dies every 34 seconds in the United States from cardiovascular disease.1
About 697,000 people in the United States died from heart disease in 2020—that’s 1 in every 5 deaths.1,2
Heart disease cost the United States about $229 billion each year from 2017 to 2018.3 This includes the cost of health care services, medicines, and lost productivity due to death.
http://www.cdc.gov/dhdsp/data_statistics/fact_sheets/fs_heart_failure.htm

## Slide 4: Disease Progression
Jessup, NEJM 2003

## Slide 5: The CardioMEMS™ HF System Provides The Actionable Data You Need To Empower Intelligent Remote Monitoring
The PA sensor is inserted via right heart catheterization.
Patient takes daily sensor reading from the comfort of their home.
Data wirelessly transmitted to Merlin.net™ PCN, a secure website that easily presents PA pressure data to inform proactive treatment modifications.
Clinician reviews data and contacts patient, as necessary.

## Slide 6: Hospitalization reduction
Chart: 
| Category 1
1-yr Pre-implant | 1.25
1-yr Post-implant | 0.54
2-yr Post-implant | 0.37
heart failure
U.S. Post-approval Study9,13
57%
Reduction
at 1 Year
0.43 (0.39, 0.47)
P < 0.0001
Supporting Studies | Heart Failure Hospitalization Reduction | Design
GUIDE-HF11 | 28% | RCT
CHAMPION1-2 | 33% | RCT
U.S. Post-approval9 | 57% | Single-arm
U.S. Post-approval 2-Year Outcomes13 | 70% | Single-arm
MEMS-HF8 | 62% | Single-arm
COAST-UK10 | 82% | Single-arm
Economic Impact6 | 34% | Retrospective Database
Propensity-matched Cohort7 | 24% | Retrospective Database
1.25
0.54
0.37
70%
Reduction
at 2 Years
0.30 (0.25, 0.35)
P < 0.0001

## Slide 7: 2021 Demographics
Financial Class | MCS Patients | IU Health
Managed Care | 54% | 53%
Medicaid/HIP | 17% | 23%
Medicare | 28% | 24%
Other/Special Contract (VA) | 1% | 0%
Uninsured | 0% | 0%
Race | MCS Patients | Indiana
White | 84% | 83%
African American | 15% | 9%
Asian | 0% | 2%
Other | 1% | 6%
Gender | MCS Patients | Indiana
Female | 27% | 51%
Male | 73% | 49%
Age | MCS Patients | Indiana
20-44 | 21% | 37%
45-64 | 46% | 18%
65-older | 33% | 15%
    Data Source: MCCM, Medical Records, and STATS Indiana

## Slide 8
2
1
1
3
1
1
ILLINOIS
Ohio
3
15
2
1
1
2
2
1
2
1
2
1
1
1
3
2
2
1
1
2
1
1
1
1
3
1
Current Patients
2022

## Slide 9: Not everyone can get a VAD!!
Clinical inclusion/exclusion criteria
  BTT
  DT
Psychosocial Evaluation
Pre-VAD meeting with the VAD team
  Dependable caregiver support
Presented to Advance Heart Care Board for acceptance for BTT/DT VAD
Advanced Therapy Evaluation
Notes: 9
9
Many things are evaluated before a patient is able to get a VAD and it is put in front of the Advanced Heart Failure Board. At times the patient is placed on a contract to meet requirements for transplant Letting the potential patient meet an existing VAD patient. Getting acceptance from the board allows a fairness for the program

## Slide 10
A MCSD is designed to circulate the blood
in the body when the patient’s heart is
unable.
Types of Mechanical Circulatory Support Devices
  HeartMate 2/3
  HeartWare
Defining a Mechanical Circulatory
Support Device (MCSD)
Notes: 10
I will begin this lecture by discussing
670,000 new cases of heart failure are diagnosed in the US every year
277,000 deaths are caused by heart failure each year
Approx 2200 hundred transplants have been done a year nation wide- number has stayed constant for the past eleven years
Approx 3000 people are waiting, the wait time can be months to years
Destination/Bridge- we have 60 patients at home

## Slide 11: 1st Generation Pulsatile LVADs
Top: Heartmate XVE
Below: PVAD
Top: Novacor
Bottom: IVAD

## Slide 12
HeartMate 2
HeartMate 3
HeartWare
Types of MCSD’s
Notes: 12
Preload and afterload, early notification is best

## Slide 13: Comparison of Pulsatile and Continuous Flow
(Both have average flow between 4-5 L/min)
Pulsatile
Continuous

## Slide 14
Left ventricular assist device only
Long term use
Axial flow – continuous flow
Bridge to transplant
Destination therapy
HeartMate 2
Notes: 14
(BSA >1.5 4.8ft and 132 lbs )(BSA >1.2 5ft 2 inches and 96lbs)

## Slide 15: HeartMate 3
Left ventricular assist device only
Long term use
Centrifugal pump - continuous flow
Bridge to transplant
Destination therapy

## Slide 16: HeartMate System Controller Alarm Indicators
Low Battery Advisory
Low Battery Hazard
Advisory Alarm
Hazard Alarm
Driveline Disconnected Hazard Alarm
Power Cable Disconnected Advisory
The system controller will alarm if it detects an issue in the operation of the pump, if the patient is experiencing alarms, contact the Implant Center for LVAS-specific emergency instructions
CAUTION – Investigational device. Limited by US Federal law to investigational
ClinicalTrials.gov Identifier: NCT02224755
10001926

## Slide 17: HeartMate 3 System Controller Viewing Pump and System Parameters
Press the Display Button to view pump parameters and backup battery charge status on the display screen.
*
Notes: IFU pg 53: Viewing information about the pump is useful when recording daily values or trying to resolve system problems on the telephone. When the System Controller is running, the user interface can display the following information about current system operations:
Speed in rpm
Mode (indicated as Pulse Mode by triangle symbol)
Flow in liters per minute
Pulsatility Index (abbreviated as “PI” on the screen)
Power in watts
Charge status of the System Controller’s 11 Volt Lithium-Ion backup battery
To view information on the user interface screen, press and release the display button. Each push of the display button brings up a new screen. Each screen illuminates for 15 seconds before it goes black, unless another button is pushed. The screens are always displayed in the same order, starting with the first (Speed) screen. A dot at the bottom of each screen provides navigational information about which of the five screens is in view.

## Slide 18: Pump Parameters
  Speed
  Flow
  Power
  Pulsatility Index
IMPORTANT!
One single pump parameter is not a surrogate for monitoring the overall clinical status of the patient. Any change in parameters should be evaluated with all clinical considerations in mind.
HeartMate 3™ LEFT VENTRICULAR ASSIST SYSTEM THEORY OF OPERATION

## Slide 19: Low Flow (Hazard Alarm)
Alarm Means:
Pump flow is less than 2.5 lpm
Constant audio tone
Silence is 2 minutes
Action:
Connect to a working power source
Check connections (Driveline, Modular cable connection, Power, and AC), and push Controller buttons
Evaluate the patient for causes
Notes: IFU pg 385
Behavior and appearance:
Flashing Red Heart ( ) on the user interface.
“Call Hospital Contact” and “Low Flow” alternate on the screen.
Alarm tone: Constant tone.
Alarm silence: 2 minutes or until a new hazard alarm occurs
System Monitor alarm active: LOW FLOW
Potential causes
Decreased preload (right heart failure, tamponade, hypovolemia, bleeding, etc.)
Obstruction of pump inflow or outflow
Systemic hypertension
Pump off or driveline disconnected
To resolve alarm:
Ensure the Driveline is connected to the System Controller.
Ensure that a power source is connected to the System Controller.
Clinically evaluate patient – ECHO to assess RV, LV function, inlet cannula obstruction
Treat underlying cause

## Slide 20: Driveline Disconnected (Hazard Alarm)
Alarm Means:
The Driveline is disconnected from the System Controller
Constant audio tone
Silence is 2 minutes
Action:
Immediately connect the driveline to the System Controller
Check Modular cable connector for a complete connection
Exchange controller
Important! If controller stops alarming Connect Driveline, but Low Speed and Low Flow alarms persist and the pump does not ramp up to the set speed, disconnect and reconnect the driveline.
Notes: IFU pg 384
Behavior and appearance
Flashing Red Heart on the user interface.
Flashing red light near Driveline connector.
“Connect Driveline” flashes on the screen.
Green “pump running” symbol is off.
Alarm tone: Constant tone.
Alarm silence: 2 minutes or until a new hazard alarm occurs.
System Monitor alarm active: DRIVELINE DISCONNECTED
Action
Immediately reconnect the Driveline to System Controller and move the Driveline Safety Lock on the System Controller to the locked position.
If alarm persists after reconnecting the Driveline, press any button on the System Controller to attempt pump start. Also, check that the Modular In-Line Connector is secure. Otherwise, check if the fixed speed setting is below 4000 rpm AND the System Controller’s backup battery is not installed. Under these conditions, the pump can only be started from the System Monitor’s Clinical or Settings screen by pressing the Pump Start button
If Driveline is connected and alarm persists, replace System Controller with a backup System Controller.

## Slide 21: No External Power (Hazard Alarm)
Alarm Means:
The System Controller is not receiving power from either power cable. The pump is being powered by the System Controller’s 11 Volt Lithium-Ion backup battery
Constant audio tone
Silence is 2 minutes
Action:
Immediately connect both power cables to a working power source (14 V Li-Ion battery, Power Module, or Mobile Power Unit)
Notes: IFU pg 385
Behavior and appearance:
Flashing Red Battery on the user interface.
“Connect Power Immediately” and Backup Battery graphic alternate on the screen.
Yellow light near the black power cable connector is flashing.
Yellow light near the white power cable connector is flashing.
Alarm tone: Constant tone.
Alarm silence: 2 minutes or until a new hazard alarm occurs.
System Monitor alarm active: NO EXTERNAL POWER
To resolve alarm:
Immediately connect to a functioning Power Module, Mobile Power Unit, or two fully-charged HeartMate 14 Volt Lithium-Ion batteries to ensure the pump does not stop.
The 11 Volt Lithium-Ion backup battery inside the System Controller provides power to the pump for at least 15 minutes when fully charged if the main power source is disconnected or fails.
IMPORTANT! If external power is not restored, the pump gradually slows to the low speed limit to save power in an effort to prevent the pump from stopping. When adequate power is supplied, the pump reverts to the previous speed and the red battery alarm clears.

## Slide 22
Left ventricular assist device only
Long term use
Centrifugal pump - continuous flow
Bridge to transplant
Destination therapy
HeartWare

## Slide 23: HeartWare® Controller Display
Scroll Button
Alarm Mute Button
AC/DC Indicator
Battery Indicator 1
Alarm Indicator
Battery Indicator 2
Controller Display
Power Source 2
Power Source 1

## Slide 24: HVAD® Controller: Alarm Indicator & Mute Button
Alarm Indicator Symbol
Low Priority: Solid Yellow
Medium Priority: Flashing Yellow
High Priority: Flashing Red
*as shown
Alarm Mute Button

## Slide 25: Alarm Guide

## Slide 26
Place the blood pressure cuff on the upper arm and apply transmission gel to the skin over the brachial artery
Position the probe over the artery until pulse is heard. Keep the pressure light enough so that the artery is not compressed
Inflate the blood pressure cuff until the arterial pulse is no longer audible
Slowly deflate the cuff until the first sound can be heard
Document this single number as the blood pressure
    Recommended range for HeartMate and Heartware patients is 60-90 mmHg, not to exceed 90 mmHg
    Review patient’s history log to assure Doppler pressure is not >90 mmHg
It is recommended to use a Doppler to obtain an accurate blood pessure
Blood Pressure Assessment1
Continuous flow device
1. Clinical Management of Continuous-flow LVADs in Heart Failure. JHLT. 2010:1-39.
10001926

## Slide 27: Physiologic Management
Optimize preload and afterload and you will optimize the pump output

## Slide 28: Physiologic Effects of Continuous Flow VADs
Continuous flow means weak or no pulse despite adequate blood supply.
Blood pressure is nonpulsatile: use MAP 60-90
Cardiac rhythm may not correlate with physiologic presentation: AFib VT
Many of these patients will have a pacemaker/ICD.
Anticoagulation increases risk of bleeding complications INR 2.0-3.0(trauma, head injury, GI)
VAD patients require significant hydration - 2 Liters/daily

## Slide 29: Emergency Situations
Assess patient and VAD monitor
OK to intubate, defibrillate, and give medications
If need to defibrillate NO disconnection required
Can DO chest compressions if you can’t get pump running
Notify Cardiologist (317)-962-2800

## Slide 30
The driveline will exit on the left or right side of the abdomen, where ingrowth will create a seal around the tubing
The aseptic dressing change is performed by the care provider according to the implanting center policy
The driveline should always be anchored, as breaking the seal will increase the chance of infection
Report any evidence of redness to the VAD Team, as they will likely treat with antibiotics
We do NOT like exercises where they utilize their abdomen muscles we have seen this cause driveline infection
Driveline Care
1. Clinical Management of Continuous-flow LVADs in Heart Failure. JHLT. 2010:1-39.
10001926
Notes: US-HM2-04150208(1)

## Slide 31: What should the patient bring to rehab???
Caregivers are excellent resources
Bring their back up bag – spare controller, batteries, troubleshooting guide
Bring their log files to review their trends at home
Consider bringing their wall adapter with heartware but not required

## Slide 32: Concerns for patients during Cardiac Rehab
Hypovolemia “Suckdown”
Blood pressure – maintaining MAP <90 but assessing your patient
Heart Rate – majority of patients have ICD
Right Heart Failure (often rhythm-related)
Bleeding issues related to anticoagulation therapy or from the pump itself; AVMs in their GI system, or stroke
Infection (may be unrelated to device)
Connections – Power and Driveline
Pump Complications (pump clot, driveline fracture, cable damage)

## Slide 33: Webinar Education
Training Easily Accessible
  We could initiate webinar for your area and record to forward to entire staff
  Ability to participate with any computer. “Show & tell with the devices”
  Case studies shared for review and discussion
Post-Presentation Resources
  PowerPoint and EMS field guide available
  Educational presentations are recorded and available via email link for those who could not join live

## Slide 34: ER and EMS Training
2013: Initiated statewide EMS guide for all MCS implanting centers including contact numbers for each facility to streamline care
Notify local ER and EMS
Quarterly webinar training with ER and EMS prior to discharge with new implant
Continue to develop a distribution list of EMS facilities and ERs as a resource to educate the community

## Slide 35: COVID and VADs
COVID vaccinations
If patients get COVID we want to try to get them IV infusion
Majority of our patients have had COVID
Encourage patients to get O2 oximeter
Stay hydrated
Frequent communication
We have NOT seen workups for LVAD or advance therapy due to post COVID heart disease

## Slide 36: Case Study 1
Mr. Smith 68 yo is a new HeartMate 3 implant discharged from implant 6 weeks. No complications with implant. Patient comes in using a rollator. Mr Smith shows up with his caregiver for cardiac rehab. Patient brings his back up bag and log files from home.
VAD Parameters at home Speed: 5500 Flow 4-4.5 PI 2.3-3.3 Power 3.4-4.0 MAPs 78-88
Patient has been having cardiac rehab for a week with no issues in vital signs
Today prior to starting rehab patients Speed 5500 Flow 4.2 PI 4.0 Power 3.6 MAP 90. During rehab MAP 95-105. VAD numbers Speed Flow 4.1 PI 4.4 Power 3.8. Patient states he feels fine. You call the VAD team
VAD Team reviews all of his numbers, When did he take his BP medication and reviewed his medication? Did he eat and drink prior to going to rehab?
PLAN:

## Slide 37: Case Study 2
Miss Jones received a HeartMate 3 and has been home for 8 weeks and going to rehab for 2 weeks. She arrives alone with her black bag. No log files
VAD Parameters today Speed 5200 Flow 4.0 PI 3.0 Power 3.6 MAP 78. She does not take her MAP at home but feels her VAD numbers are what they are at home.
During rehab MAP 95-105. VAD numbers Speed Flow 2.9 PI 8.4 Power 3.8. Patient states she doesn’t feels well. You call the VAD team
VAD Team reviews all of her numbers, When did she take her BP medication and reviewed her medication list? Did she eat and drink prior to going to rehab? Patient is on metoprolol 50 mg bid which she did not take prior to coming to rehab. She did not have any breakfast before rehab.
PLAN:

## Slide 38: Questions

## Slide 39: Thank You